## Slide 1: Flow Cytometry.
Malcolm Needs CSci FIBMS FBBTS

## Slide 2: This lecture was kindly written by Gordon Burgess.
Malcolm Needs CSci FIBMS FBBTS

## Slide 3: This lecture is designed to give an overview of:
Definitions.
Basics of flow cytometry.
Fluidics.
Malcolm Needs CSci FIBMS FBBTS

## Slide 4: This lecture is designed to give an overview of:
Various scatters.
Detectors and filters.
Malcolm Needs CSci FIBMS FBBTS

## Slide 5
Definitions.
Measuring properties of cells in flow.
Flow sorting.
Sorting (separating) cells based on properties measured in flow.
Also called fluorescence-activated cell sorting (FACS).
Malcolm Needs CSci FIBMS FBBTS

## Slide 6
Basics of Flow Cytometry.
Cells in suspension
Flow in single-file through
An illuminated volume where they
Scatter light and emit fluorescence
That is collected, filtered and
Converted to digital values
That are stored on a computer.
Fluidics
Optics
Electronics
Malcolm Needs CSci FIBMS FBBTS

## Slide 7
Fluidics.
Need to have cells in suspension flow in single file through an illuminated volume.
In most instruments, this is accomplished by injecting a sample into a sheath fluid as it passes through a small (50-300 µm) orifice.
Malcolm Needs CSci FIBMS FBBTS

## Slide 8
Fluidics.
When conditions are right, sample fluid flows in a central core that does not mix with the sheath fluid. This is termed “laminar flow”.
The introduction of a large volume into a small volume, in such a way that it becomes “focused” along an axis, is called “hydrodynamic focusing”.
Malcolm Needs CSci FIBMS FBBTS

## Slide 9
Fluidics.
Notice how the ink is focused into a tight stream as it is drawn into the tube under laminar flow conditions.
Notice also how the position of the inner ink stream is influenced by the position of the ink source.
Malcolm Needs CSci FIBMS FBBTS

## Slide 10
Flow Cell.
Injector
tip
Fluorescence
signals
Focused laser
beam
Sheath
fluid
Malcolm Needs CSci FIBMS FBBTS

## Slide 11
Optics - Forward Scatter Channel.
When a laser light source is used, the amount of light scattered in the forward direction (along the same axis that the laser light is travelling) is detected in the forward scatter channel.
The intensity of forward scatter is proportional to the size, shape and optical homogeneity of cells (or other particles).
Malcolm Needs CSci FIBMS FBBTS

## Slide 12
Forward Scatter Channel.
FSC Sensor
Laser
Malcolm Needs CSci FIBMS FBBTS

## Slide 13
Optics - Side Scatter Channel.
When a laser light source is used, the amount of light scattered to the side (perpendicular to the axis that the laser light is travelling) is detected in the side or 90o scatter channel.
The intensity of side scatter is proportional to the size, shape and optical homogeneity of cells (or other particles).
Malcolm Needs CSci FIBMS FBBTS

## Slide 14
90o Light Scatter.
FSC Sensor
SSC Sensor
Laser
Malcolm Needs CSci FIBMS FBBTS

## Slide 15
Optics - Light Scatter.
Forward scatter tends to be more sensitive to surface properties of particles (e.g. cell ruffling) than side scatter.
  Can be used to distinguish live from dead cells.
Malcolm Needs CSci FIBMS FBBTS

## Slide 16
Side scatter tends to be more sensitive to inclusions within cells than forward scatter.
  Can be used to distinguish granulated cells from non-granulated cells.
Malcolm Needs CSci FIBMS FBBTS

## Slide 17
Light Scatter Gating.
90 Degree Scatter
0
200
400
600
800
1000
8
15
20
30
40
50
100
200
1000
Lymphocytes
Monocytes
Neutrophils
Side Scatter Projection
Scale
Malcolm Needs CSci FIBMS FBBTS

## Slide 18
Optics - Fluorescence Channels.
The fluorescence emitted by each fluorochrome is usually detected in a unique fluorescence channel.
The specificity of detection is controlled by the wavelength selectivity of optical filters and mirrors.
Malcolm Needs CSci FIBMS FBBTS

## Slide 19
Fluorescence Detectors.
Laser
Fluorescence
FSC Sensor
Fluorescence detector
(PMT3, PMT4 etc.)
Malcolm Needs CSci FIBMS FBBTS

## Slide 20
Optics - Filter Properties.
Optical filters are constructed from materials that absorb certain wavelengths (while transmitting others).
Transitions between absorbance and transmission are not perfect; The sharpness can be specified during filter design.
Malcolm Needs CSci FIBMS FBBTS

## Slide 21
Optics - Filter Properties.
Long pass filters transmit wavelengths above a cut-on wavelength.
Short pass filters transmit wavelengths below a cut-off wavelength.
Band pass filters transmit wavelengths in a narrow range around a specified wavelength.
  can be specified Band width.
Malcolm Needs CSci FIBMS FBBTS

## Slide 22: Standard Long Pass Filters.
Transmitted light
Light source
520 nm long pass filter
>520 nm light
Transmitted light
Light source
575 nm short pass filter
<575 nm light
Standard Short Pass Filters.
Malcolm Needs CSci FIBMS FBBTS

## Slide 23
Standard Band Pass Filters.
Transmitted light
White light source
630 nm band pass filter
620 - 640 nm light
Malcolm Needs CSci FIBMS FBBTS

## Slide 24
Optics - Filter Properties.
When a filter is placed at a 45o angle to a light source, light that would have been transmitted by that filter is still transmitted but light that would have been blocked is reflected (at a 90o angle).
Used this way, a filter is called a dichroic filter or dichroic mirror.
Malcolm Needs CSci FIBMS FBBTS

## Slide 25: Optical Filters.
Dichroic filter/mirror at 45o.
Reflected light
Transmitted light
Light source
Malcolm Needs CSci FIBMS FBBTS

## Slide 26
Bypass filters
Example channel layout for laser-based flow cytometry.
Laser
Flow cell
PMT 1
PMT 2
PMT 3
PMT 4
Dichroic filters
Malcolm Needs CSci FIBMS FBBTS

## Slide 27
Ethidium
PE
cis-Parinaric acid
Texas red
PE-TR conj.
PI
FITC
600 nm
300 nm
500 nm
700 nm
400 nm
457
350
514
610
632
488
Common
laser
lines
Malcolm Needs CSci FIBMS FBBTS

## Slide 28: Propidium Iodide.
400 nm
500 nm
600 nm
700 nm
PI
DNA
Excitation
Emission
300 nm 400 nm 500 nm 600 nm 700 nm
Malcolm Needs CSci FIBMS FBBTS

## Slide 29: Fluorescein (FITC).
400 nm
500 nm
600 nm
700 nm
Wavelength
Protein
Excitation
Emission
300 nm 400 nm 500 nm 600 nm 700 nm
Malcolm Needs CSci FIBMS FBBTS

## Slide 30: Phycoerytherin (PE).
Excitation
Emission
300 nm 400 nm 500 nm 600 nm 700 nm
Protein
Malcolm Needs CSci FIBMS FBBTS

## Slide 31: Allophycocyanin (APC).
Protein
632.5 nm (HeNe)
Excitation
Emission
300 nm 400 nm 500 nm 600 nm 700 nm
Malcolm Needs CSci FIBMS FBBTS

## Slide 32
Optics – Detectors.
Two common detector types.
  Photodiode
    Used for strong signals when saturation is a potential problem (e.g., forward scatter detector).
  Photomultiplier tube (PMT)
    More sensitive than photodiode, but can be destroyed by exposure to too much light.
Malcolm Needs CSci FIBMS FBBTS

## Slide 33
488 nm laser
+
-
Fluorescence Activated
Cell
Sorting
Charged plates
Single cells sorted
into test tubes
FSC Sensor
Fluorescence detector
Malcolm Needs CSci FIBMS FBBTS